## Standard Operating Procedure

## Large Laboratory Equipment Decontamination

## Protocol

- Identify contamination: What material has this equipment been used with:

        - Chemical: (Contact EH&S lab safety for assistance)

        - Biological: (Contact biosafety@ehs.ucsb.edu for consultation)

        - Radiological: (Contact EH&S Radiation Safety for assistance)

[NOTE: Perchloric acid can build up in equipment and generate shock sensitive metal perchlorates. If perchloric acid contamination is possible, contact EH&S lab Safety.]

- Identify appropriate cleaning/neutralizing/disinfecting agents for the identified contaminants.

- De-energize system (electrical, mechanical) using appropriate procedures. Refer to manufacturer’s instructions as needed. Use lock-out/tag-out protocols if the

decontamination takes multiple days or if the equipment will be otherwise unattended for a period of time.

- Most equipment used with biological equipment may be decontaminated by wiping down surfaces with a high-level disinfectant.

  - Defrost freezers prior to surface decontamination.

  - Biosafety cabinets surfaces are decontaminated before a vapor phase decontamination process by an external vendor. Biosafety cabinets should be decontaminated with a vapor phase process before relocating, moving or re-allocating to new faculty.

  - Secondary fungal growth in flasks used for liquid waste deactivation is treated with additional full strength household bleach.

  - Removable incubator parts (shelves, racks, humidity pan) may be autoclaved and then wiped down with 70% ethanol.

- Drain fluids, if necessary, using appropriate procedures (pump/pneumatic oils, refrigerant). Refer to manufacturer’s instructions as needed.

- Confirm that any storage compartments in the equipment are empty.

- Cleaning Procedure:

  - Avoid dry brushing, wiping, blowing with compressed air, or vacuuming as this can aerosolize contaminants and create a significant respiratory hazard.

  - Moisten a paper towel with the appropriate cleaning agent, and wipe all surfaces gently, to avoid kicking up particulate matter. Progress from less dirty areas to more dirty areas. Dispose of soiled towels in a solid waste bag.

A second pass with the moistened towels may be necessary to decontaminate some surfaces. As the first pass will have removed most loose powder, a more aggressive scrubbing is now possible.

- If removing radiological contamination, contact the Radiation Safety group at Environmental Health & Safety to confirm successful decontamination.

- Remove any chemical, biohazard or radiological warning labels or signage.

- If the equipment is to be disposed of, relocated, repaired or put into storage, please fill out the attached Laboratory Equipment Decontamination Form.

## Personal Protective Equipment

- Respiratory Protection

As large equipment is often too large to move into a fume hood or other local exhaust device, the inhalation of dusts and fumes is of primary concern. All respiratory protection must be obtained through the campus Respiratory Protection Program.   Please contact this group to get advice on and access to the appropriate equipment.

- Eye Protection

At a minimum, ANSI compliant safety glasses with side shields must be worn. Chemical splash goggles should be worn if working with larger quantities of liquid, or if a splash risk is present. If the cleaning agent or the contaminants are skin-contact hazards, a face shield should be worn.

- Skin and body protection

A well-fitted chemically resistant lab coat must be worn, along with long pants and closed-toe shoes. A flame-resistant lab coat is required when working with flammable or pyrophoric liquids.

A chemical resistant apron should be worn when using larger volumes of liquid, or if a splash risk is present.

A well-fitted liquid impervious lab coat must be worn, along with long pants and closed-toe shoes, while decontaminating biological materials from equipment.

- Hand protection

At minimum, disposable nitrile gloves must be worn. Consult the manufacturer's glove compatibility chart to ensure that the glove you plan on using are compatible with the cleaning agents and the contaminants.

Additional personal protective equipment may be required if the specific procedure or contaminants present additional risk. It is the PI/Supervisor’s responsibility to ensure that any additional PPE requirements are identified and communicated to researchers. Contact EH&S for consultation.

## Waste Management

- All chemically contaminated solids (e.g. wipes, paper towels, etc.) should be bagged, labeled with the appropriate hazards and disposed of as solid hazardous waste through the EH&S hazardous Waste Program.

- All liquid chemical waste should be stored in a compatible sealable container, labeled with the appropriate hazards and disposed of as liquid hazardous waste through the EH&S hazardous Waste Program.

- Bleach solutions may be dispensed down the sink with an excess of water.

Radiologically contaminated waste should be treated and/or disposed of via the requirements of the Radiation Safety group.

Laboratory Equipment Decontamination Form

This equipment is being decontaminated for: Disposal ☐  Relocation ☐  Repair ☐  Storage ☐

Type of Equipment/Model:		Serial #:

Current location (building/room):

Department:	Owner:			Phone:

Address the following issues as needed for decontamination:

☐  Stored Energy (e.g. electrical, pneumatic) discharged or depressurized			☐NA

☐  Fluids drained and appropriately captured or disposed of (e.g. pump oil, refrigerant)		☐NA

☐  Radioactive Materials used in the equipment							☐NA

☐  Fully monitored for radioactive materials, has been decontaminated, and is not radioactive

☐  Radiation hazard label removed

☐  EH&S Radiation Safety Officer has cleared the equipment

RSO Signature				Name					Date

☐  Hazardous Chemicals used in the equipment					☐NA

☐  Removed, cleaned and/or neutralized with appropriate cleaning agent

☐Chemical warning labels removed or defaced

☐  Biological Materials used in the equipment						☐NA

☐Deactivated with high level disinfectant

☐Biohazard labels removed or defaced

I certify that, to the best of my knowledge, the equipment is free of hazardous materials or hazards, including those noted above.

Signature:				Name:					Date: