Electron Microscope Radiation Safety Guide

Radiation safety concerns regarding electron microscopes are minimal, and if anything, X-ray radiation is only produced from the backscattered electrons impinging on the sample. Electron microscopes are well shielded, including viewing ports with lead impregnated glass, such that any X-rays generated internally do not penetrate outside the unit.  However, the California Department of Public Health requires these units to be registered with the state and that operators have a minimal awareness of radiation safety related to machine operation. Contact EH&S, Radiation Safety Office, when acquiring or disposing of electron microscopes.

- Operators must be familiar with the operation of the unit and must be authorized by the laboratory supervisor to use the microscope. A copy of the microscope operating instructions needs to be readily available and this guideline incorporated into the operator training.

- The manufacturer’s original unit design must not be modified (e.g., viewing ports removed, etc), which may compromise the radiation shielding of the unit.

- A radiation survey for X-ray leakage will be performed:

    - Initially at time of installation

    - Whenever the microscope is modified or attachments/accessories are added/removed, which may compromise the radiation shielding

    - If possible, radiation surveys should be made under "worse case"  conditions, such as:

-highest acceleration voltage and beam current

-low magnification

-beam at crossover on specimen

-all apertures removed

- A log book should be kept of any maintenance done on the machine. The Radiation Safety Office, EH&S (ext. 3588 or 7256) must be notified if any modifications are made to the interlocks or any other safety devices.

- EH&S, Radiation Safety, will periodically conduct radiation surveys and verify the awareness of the operator to radiation safety of machine use.

- Radiation levels should be less that 0.5 mR at 5 centimeters from the unit (measured by Radiation Safety, EH&S). Essentially, all electron microscopes will have radiation readings which are indistinguishable from normal background radiation.